## Slide 1
Jannik Ebert, Miguel-David Mendez-Bohorquez, Robert Schmoll, Andreas Kroll
QIRT 2026
Naples, Italy, July 1st 2026
Modeling the Size-of-Source Effect in Thermography
Using a Measured MTF Extended by a Parametric Scattering Model

## Slide 2: Motivation
Size-of-Source Effect
Input Signals
Output Signals
Modulation Transfer Function
Aim: Measure the MTF and predict the SSE
System
Problem: MTF is based on diffraction in the optics and affects only small objects
Solution: Extend the MTF with a Parametric Model to include additional effects
SSE
(calculated from MTF)
Modeled SSE
Measured temperature
depends on object size
Diameter in px
Temperature in °C
Notes: Emissivity
Diffraction based MTF  Deshalb unvollständige Beschreibung SSE

## Slide 3: Outline
MTF fundamentals
Influence of scattering on the MTF
Modeling approach
Experimental setup
Comparison between measured and modeled SSE

## Slide 4: Modulation Transfer Function (MTF)
MTF = ourier transform of the spatial impulse response (PSF)  High spatial frequency components are suppressed
*
=
MTF
Spatial frequency
100%
PSF
INPUT
OUTPUT
Modulation
diffraction limit
high frequency signals (small objects, fine details) can‘t be resolved
relevant for large objects in the image
large measurement deviations
Notes: MTF linear 1D!

## Slide 5: MTF Measurement and Extension
Evaluation of the MTF with the edge response (BS ISO 12233:2017)
sharp peak: high spatial frequency contributions
wings:
low spatial frequency contributions (impractical to measure)
fraction of radiation contained in the wings
~ width of the wings
measured
parametric
model the wings in the spatial frequency domain (inspired by the Harvey-Shack theory)
diffraction based
Slanted edge
edge spread function
line spread function

## Slide 6: Modeling Approach
multiplication
with MTF and STF
Diameter
Calculating the SSE with the measured MTF and the parametric STF
measured value:
proposed definition:
„contrast amplification“ of the input signal
1
SSE
0

## Slide 7: Experimental Setup (MTF Measurement)
Optris PI450
(LWIR)
coated metal
sheet
heat plate radiator
collect one thermogram of the slanted edge (contains 50 frames to account for noise)
Notes: SSE: time sequence

## Slide 8: Experimental Setup (SSE Measurement)
Identical setup to MTF measurement
 replace slanted edge with iris diaphragms
circle detection
(position and diameter)
(average radiance)
Open and close the iris diaphragms for 2 cycles  times series
FOV
FOV
close
close
open
open

## Slide 9: Parameter optimization
Cost function:
Optimization strategy:
Differential evolution
initial estimate
Nelder-Mead algorithm
 refinement

## Slide 10: Results (Modeled and Measured SSE)
Parameters found using a bootstrap analysis
with 20 realizations
(radiance values)
detector width (382 px)
Notes: Text hinzufügen (Stichpunkte, wichtigste)
Ergebnisse Tabelle

## Slide 11: Conclusion & Outlook
The proposed model
  Enables a descripton of the SSE within a linear systems framework
  Significantly improves the characterization of the SSE by incorporating additional effects
  Enables modeling of individual camera components (lenses, optical filters, …)
In future work
  Extend the model to 2D
  Derive a convolutional filter to compensate the SSE  reduce measurement uncertainty
  Inform the model with physical kowledge

## Slide 12: End of the Presentation
Thank you for your attention!
Contact details:
Jannik Ebert (B. Sc.)
Measurement and Control Research Group
University of Kassel, Kassel, Germany
mail: jannik.ebert@mrt.uni-kassel.de
visit us: www.uni-kassel.de/fb15/mrt

## Slide 13: Influence of the parameters on the modeled SSE
SSE becomes more profound
SSE converges more slowly
increases:
increases:

## Slide 14: Results (Modeled and Measured MTF)
Notes: Text hinzufügen

## Slide 15: Camera Specifications

## Slide 16: Edge-Spread-Function Algorithm